## Shutdown FELIX quick guide

@ FELIX console

- Turn off the GUN RF TIMING ON, 3GHZ RF TIMING ON

- If FELICE is not running turn off the GUN HV ON, GUN RF ENABLE, 3 GHZ ENABLE, CHARGE MODULATOR and SLINGERCIRCUIT ON

- Close vacuum line 2B

- Put the timer FELIX BEAM PULSE to 2µs & Repetition rate to 1 Hz

- If FELICE is still running turn of timer S2 FELIX accel SP 2

- Disable low level RF Feedback units

- FEL2: switch magnet 2B20_25 to Main power off

If FELICE is running skip step 8

- Close all vacuum valves

If FLARE or FELICE is still running proceed to step 11

@ Building console

- Turn MASTERKEY from RUN to SWEEP

@ FLARE console

- Turn off the solenoids from FLARE

@ FEL Technical Area

- Turn off the Klystron WEST Amplifier.

- Put the main Modulator WEST cabinet in LOCAL control and turn off   high voltage, line, klystron and cooling

- Turn off the High voltage power supplies from both modulators!

- Put the main Modulator EAST cabinet in LOCAL control and turn off   high voltage, line, klystron and cooling

- Turn off the gun RF amplifier and power supply.

- Turn off the prebuncher power supply.

- Turn off all the cabinets with the green buttons.